LabMonitor

Lab Environment Monitoring System

Technical Documentation & Setup Guide

Deeplight • 9 March 2026

# Table of Contents

# 1. Overview

LabMonitor is a lightweight data-acquisition and visualisation stack that continuously monitors the physical environment of a laboratory. It collects temperature and humidity data from a Thorlabs TSP01B USB sensor, fetches outdoor weather data from OpenWeatherMap, persists everything in InfluxDB 2.x, and exposes live dashboards in Grafana.

## 1.1 Architecture

| Component | Technology | Role |
| --- | --- | --- |
| Sensor driver | Python + ctypes / TLTSPB_64.dll | Reads temperature & humidity from TSP01B over USB |
| TSP01 logger | tsp01_logger.py | Polls sensor every 10 s, writes to InfluxDB |
| Weather logger | owm_logger.py | Polls OpenWeatherMap every 10 min, writes to InfluxDB |
| Time-series DB | InfluxDB 2.x (Docker) | Stores all measurements in bucket LabEnvironment |
| Visualisation | Grafana (Windows service) | Dashboards accessible on LAN at port 3000 |
| Task scheduler | Windows Task Scheduler | Starts both loggers at system boot, no console window |

## 1.2 Data Flow

TSP01B sensor  →  tsp01_logger.py  →  InfluxDB (measurement: tsp01)  →  Grafana

OpenWeatherMap API  →  owm_logger.py  →  InfluxDB (measurement: openweathermap)  →  Grafana

## 1.3 Measurements Stored

| Measurement | Tags | Fields |
| --- | --- | --- |
| tsp01 | location, sensor_id | temperature_internal_c, humidity_pct, temperature_external_c (optional) |
| openweathermap | name, sys_country | main_temp, main_humidity, main_pressure, wind_speed, wind_deg, clouds_all, visibility, weather_0_description, … |

# 2. Prerequisites

## 2.1 Hardware

- Thorlabs TSP01B USB temperature & humidity sensor

- Windows PC (64-bit) with the sensor connected via USB

- Internet connection (for OpenWeatherMap)

## 2.2 Software

| Software | Version | Notes |
| --- | --- | --- |
| Thorlabs TSP01 software | latest | Installs TLTSPB_64.dll — mandatory driver |
| Python (Miniconda) | 3.10+ | C:\Users\Administrator\miniconda3\ |
| Docker Desktop | latest | Runs InfluxDB container |
| Grafana | 10+ (Windows service) | Installed at C:\Program Files\GrafanaLabs\grafana\ |
| NSSM | 2.24+ | Bundled with Grafana installer at svc-12.x.x.x\nssm.exe |

## 2.3 External Accounts

- OpenWeatherMap free account — obtain an API key at openweathermap.org/api

# 3. Installation from Scratch

## 3.1 Install Thorlabs TSP01 Software

- Download the TSP01 software package from:

https://www.thorlabs.com/software_pages/ViewSoftwarePage.cfm?Code=TSP01

- Run the installer and follow the prompts.

- Verify the driver DLL was installed:

dir "C:\Program Files\IVI Foundation\VISA\Win64\Bin\TLTSPB_64.dll"

Important: keep the Thorlabs TSP01 application closed while the logger is running. The two processes cannot share the USB device simultaneously.

## 3.2 Install Python Dependencies

- Open a Miniconda PowerShell prompt and navigate to the project folder:

cd "C:\Users\Administrator\OneDrive - Deeplight\Documents\DLT Codes\LabMonitor"

- Install required packages:

pip install -r requirements.txt

requirements.txt contains:

influxdb-client>=1.36.0 # InfluxDB 2.x Python client requests>=2.28.0 # HTTP client for OpenWeatherMap API pyyaml>=6.0 # YAML config file parsing

## 3.3 Deploy InfluxDB with Docker

- Pull and start InfluxDB:

docker run -d \ --name=LabMonitor \ -p 8086:8086 \ --restart=unless-stopped \ influxdb:latest

- Open the InfluxDB UI at http://localhost:8086 and complete the initial setup:

  - Organisation:  Deeplight

  - Bucket:        LabEnvironment

  - Note down the generated API token.

- Create a DBRP mapping so Grafana can query via InfluxQL:

$token = "<your-influxdb-token>" $org = "b375cc849f21eb88" # replace with your org ID $bucket = "33e77c0b7061778b" # replace with your bucket ID Invoke-RestMethod -Method POST \ -Uri "http://localhost:8086/api/v2/dbrps" \ -Headers @{ Authorization = "Token $token"; 'Content-Type' = 'application/json' } \ -Body ('{"bucket_id":"' + $bucket + '","database":"LabEnvironment","default":true,' + '"org_id":"' + $org + '","retention_policy":"autogen"}')

The DBRP mapping is required for InfluxQL compatibility with InfluxDB 2.x. Without it, Grafana's InfluxQL queries will not work.

## 3.4 Configure config.yaml

Edit config.yaml in the project folder to match your environment:

sensor: visa_resource: "4883_33018_M01241742" # use --list-devices to find yours timeout_ms: 5000 influxdb: url: "http://localhost:8086" token: "<your-influxdb-token>" org: "Deeplight" bucket: "LabEnvironment" measurement: "tsp01" tags: location: "lab" sensor_id: "tsp01_1" acquisition: interval_seconds: 10 retry_delay_seconds: 30 max_consecutive_errors: 5 openweathermap: api_key: "<your-owm-api-key>" city_id: 2659994 # Lausanne — find others at openweathermap.org/find units: "metric" interval_seconds: 600

To discover the TSP01 resource string (if different from above):

python tsp01_logger.py --list-devices

## 3.5 Test the Loggers Manually

- Test the TSP01 logger (should print readings every 10 s):

cd "C:\Users\Administrator\OneDrive - Deeplight\Documents\DLT Codes\LabMonitor" python tsp01_logger.py

- Test the OpenWeatherMap logger (should print weather data every 10 min):

python owm_logger.py

- Press Ctrl+C to stop. Both loggers stop within ~1 second.

- Verify data in InfluxDB UI: open http://localhost:8086, go to Data Explorer, and query the LabEnvironment bucket.

## 3.6 Register as Windows Scheduled Tasks (Background Services)

Both loggers are registered as Windows Scheduled Tasks that start at system boot and run silently (no console window) using pythonw.exe.

- TSP01 logger task:

$action = New-ScheduledTaskAction \ -Execute "C:\Users\Administrator\miniconda3\pythonw.exe" \ -Argument "tsp01_logger.py" \ -WorkingDirectory "C:\Users\Administrator\OneDrive - Deeplight\Documents\DLT Codes\LabMonitor" $trigger = New-ScheduledTaskTrigger -AtStartup $settings = New-ScheduledTaskSettingsSet -ExecutionTimeLimit ([TimeSpan]::Zero) Register-ScheduledTask -TaskName "TSP01 logger" -Action $action -Trigger $trigger -RunLevel Highest -User "Administrator" -Settings $settings -Force

- OpenWeatherMap logger task:

$action = New-ScheduledTaskAction \ -Execute "C:\Users\Administrator\miniconda3\pythonw.exe" \ -Argument "owm_logger.py" \ -WorkingDirectory "C:\Users\Administrator\OneDrive - Deeplight\Documents\DLT Codes\LabMonitor" $trigger = New-ScheduledTaskTrigger -AtStartup $settings = New-ScheduledTaskSettingsSet -ExecutionTimeLimit ([TimeSpan]::Zero) Register-ScheduledTask -TaskName "OWM logger" -Action $action -Trigger $trigger -RunLevel Highest -User "Administrator" -Settings $settings -Force

- Start both immediately without rebooting:

schtasks /run /tn "TSP01 logger" schtasks /run /tn "OWM logger"

- Verify both are running:

Get-Process pythonw

Two pythonw processes should be listed.

## 3.7 Configure Grafana

- Grafana is installed as a Windows service managed by NSSM. Verify it is running:

Get-Service Grafana

- Enable anonymous LAN access. Create the file:

C:\Program Files\GrafanaLabs\grafana\conf\grafana.ini

with the following content:

[auth.anonymous] enabled = true org_name = Deeplight org_role = Viewer

Note: If grafana.ini is not read by the service, set these values as environment variables via NSSM instead (see Troubleshooting section).

- Restart Grafana:

Stop-Service Grafana; Start-Service Grafana

- Open port 3000 in Windows Firewall:

New-NetFirewallRule -DisplayName "Grafana" -Direction Inbound -Protocol TCP -LocalPort 3000 -Action Allow

- Open Grafana in a browser at http://192.168.0.251:3000 and log in as admin.

- Add the InfluxDB datasource (Settings → Data Sources → Add):

  - Type: InfluxDB

  - Query language: InfluxQL

  - URL: http://localhost:8086

  - Authorization header: Token <your-influxdb-token>

  - Database: LabEnvironment

- Import or create dashboards. The direct URL for a dashboard is:

http://192.168.0.251:3000/d/<dashboard-uid>/<dashboard-name>

# 4. Day-to-Day Operation

## 4.1 Starting / Stopping the Loggers

| Action | Command |
| --- | --- |
| Start TSP01 logger | schtasks /run /tn "TSP01 logger" |
| Start OWM logger | schtasks /run /tn "OWM logger" |
| Stop TSP01 logger | Get-Process pythonw ¦ Stop-Process (stops all pythonw) |
| Check running loggers | Get-Process pythonw |
| Check task status | schtasks /query /tn "TSP01 logger" /fo LIST /v |

## 4.2 Log Files

Both loggers write a log file to the project directory alongside the scripts:

| File | Contents |
| --- | --- |
| tsp01_logger.log | Timestamped sensor readings, errors, and status messages |
| owm_logger.log | Timestamped weather fetches, errors, and status messages |

To watch the TSP01 log in real time from PowerShell:

Get-Content tsp01_logger.log -Wait -Tail 20

## 4.3 Configuration Reference

| Parameter | Default | Description |
| --- | --- | --- |
| acquisition.interval_seconds | 10 | How often TSP01 is polled (seconds) |
| acquisition.retry_delay_seconds | 30 | Wait between retries after a read error |
| acquisition.max_consecutive_errors | 5 | How many errors before the logger exits |
| openweathermap.interval_seconds | 600 | How often weather is fetched (seconds) |
| openweathermap.city_id | 2659994 | OpenWeatherMap city ID (2659994 = Lausanne) |
| openweathermap.units | metric | Temperature units: metric / imperial / standard |

# 5. Troubleshooting

| Symptom | Cause | Fix |
| --- | --- | --- |
| --list-devices finds no devices | TSP01 software not installed, or USB disconnected | Install Thorlabs TSP01 software; check USB cable |
| 0xBFFC04FF communication server error | Thorlabs TSP01 application is open | Close the Thorlabs TSP01 application before starting the logger |
| 0xBFFF0011 RSRC_NFOUND | visa_resource string is wrong | Run tsp01_logger.py --list-devices and copy the exact string into config.yaml |
| WritePrecision error on startup | Wrong influxdb-client version | pip install influxdb-client>=1.36.0 |
| Grafana shows no data | DBRP mapping missing, wrong token, or wrong datasource type | Verify datasource uses InfluxQL (not Flux); re-create DBRP mapping; check token |
| OWM data missing in Grafana | owm_logger not running or wrong measurement name | Check owm_logger.log; measurement must be 'openweathermap' |
| grafana.ini not read by service | NSSM does not pass a --config flag; file may not be picked up | Set environment variables via NSSM (see below) |
| Script takes long to stop | Older version without interruptible sleep | Ensure both scripts contain the _interruptible_sleep function |

## 5.1 Setting Grafana Config via NSSM Environment Variables

If grafana.ini is ignored, use NSSM to inject the settings as environment variables. Grafana maps GF_<SECTION>_<KEY> automatically:

& "C:\Program Files\GrafanaLabs\svc-12.0.2.0\nssm.exe" set Grafana AppEnvironmentExtra \ "GF_AUTH_ANONYMOUS_ENABLED=true" \ "GF_AUTH_ANONYMOUS_ORG_NAME=Deeplight" \ "GF_AUTH_ANONYMOUS_ORG_ROLE=Viewer" Stop-Service Grafana; Start-Service Grafana

# 6. Removal

Follow the steps below to completely remove LabMonitor from the machine.

## 6.1 Stop and Remove Scheduled Tasks

# Stop running logger processes Get-Process pythonw -ErrorAction SilentlyContinue ¦ Stop-Process # Delete the scheduled tasks schtasks /delete /tn "TSP01 logger" /f schtasks /delete /tn "OWM logger" /f

## 6.2 Remove the InfluxDB Docker Container and Data

# Stop and remove the container docker stop LabMonitor docker rm LabMonitor # (Optional) remove the InfluxDB image docker rmi influxdb:latest # (Optional) remove Docker volumes holding InfluxDB data docker volume ls # find volume name docker volume rm <volume-name>

Warning: removing the Docker volume permanently deletes all historical measurement data. Back up data first via the InfluxDB UI (Data → Buckets → Export) if needed.

## 6.3 Revert Grafana Anonymous Access

To re-enable the login prompt, undo the changes made in Section 3.7:

- Delete or empty the custom config file:

Remove-Item "C:\Program Files\GrafanaLabs\grafana\conf\grafana.ini" -ErrorAction SilentlyContinue

- If NSSM environment variables were set, remove them:

& "C:\Program Files\GrafanaLabs\svc-12.0.2.0\nssm.exe" set Grafana AppEnvironmentExtra ""

- Remove the firewall rule:

Remove-NetFirewallRule -DisplayName "Grafana"

- Restart Grafana:

Stop-Service Grafana; Start-Service Grafana

## 6.4 Delete Project Files

Remove-Item -Recurse -Force "C:\Users\Administrator\OneDrive - Deeplight\Documents\DLT Codes\LabMonitor"

## 6.5 Uninstall Python Packages (Optional)

pip uninstall influxdb-client requests pyyaml -y

## 6.6 Uninstall Thorlabs Driver (Optional)

Use Windows Settings → Apps → Thorlabs TSP01 to uninstall the Thorlabs software. This removes TLTSPB_64.dll and the IVI Foundation VISA driver.

# 7. File Reference

| File | Description |
| --- | --- |
| tsp01_logger.py | Main sensor acquisition daemon. Reads TSP01B via TLTSPB_64.dll and writes to InfluxDB at the configured interval. |
| owm_logger.py | OpenWeatherMap logger. Fetches current weather from the free /data/2.5/weather endpoint and writes to InfluxDB. |
| config.yaml | Shared configuration file. Contains InfluxDB credentials, sensor resource string, acquisition intervals, and OWM settings. |
| requirements.txt | Python package requirements. Install with: pip install -r requirements.txt |
| list_dll_exports.py | Developer utility. Enumerates exported function names from a DLL using a pure-Python PE parser (no external dependencies). Useful for debugging driver issues. |
| tsp01_logger.log | Runtime log for the TSP01 logger (auto-created on first run). |
| owm_logger.log | Runtime log for the OWM logger (auto-created on first run). |

# 8. Quick-Reference Card

| Task | Command / Location |
| --- | --- |
| Find TSP01 resource string | python tsp01_logger.py --list-devices |
| Test TSP01 logger | python tsp01_logger.py |
| Test OWM logger | python owm_logger.py |
| Start both loggers (no window) | schtasks /run /tn "TSP01 logger" && schtasks /run /tn "OWM logger" |
| Check loggers running | Get-Process pythonw |
| Stop all loggers | Get-Process pythonw ¦ Stop-Process |
| InfluxDB UI | http://localhost:8086 |
| Grafana (local) | http://localhost:3000 |
| Grafana (LAN) | http://192.168.0.251:3000 |
| Dashboard (OWM) | http://192.168.0.251:3000/d/7wISvn_iz/open-weather-map |
| InfluxDB Docker container | docker ps (name: LabMonitor) |
| Grafana Windows service | Get-Service Grafana |
| Project folder | C:\Users\Administrator\OneDrive - Deeplight\Documents\DLT Codes\LabMonitor |